[image: image1.png]



Nazareth, 06 December 2010.

Concerning:

State of the Chemidoc system, after fire at UCL, Tour Vésale, 3rd floor, 1200 Brussels.
On-site visit was performed by Patrick Couchement, senior service engineer on the 1st of December 2010.

The system has been taken to our workshop for investigation.

Included in the system are:
-Universal hood 2 (Sn. 76S/09436)

-Camera (Sn. A08B835008)

-PC (Sn. D24010631601230)

-Keyboard, mouse and screen

-Power- and connection cables

Findings:

-Universal Hood 2: cleaning possible.
-Camera: soot has entered the part, this is a non-serviceable part and therefore no guarantees can be given regarding the caused damage.
-PC, keyboard, mouse and screen: Not repairable due to soot damage.
Hereby I declare that the repair costs of this system will exceed the cost of a new system. 

Patrick Couchement

Senior Service Engineer

Bio-Rad Laboratories NV/SA

Begoniastraat 5

B-9810 Nazareth

